## Sheet1
Number | Category | Description | Part name | Catalog number | Vendor | Need per setup | unit | Cost per unit | Total cost | website | Notes | JAX or Collaborator Lab
1 | open field | Open field box (with front base opening for panel options) |  |  |  | 1 | [5 empty] | JAX
2 | open field | Flat front panel for open field box |  |  |  | 1 |  |  |  |  | Panel consists of 2 machined pieces of PVC sheet. | JAX
3 | open field | Front panel with food grate and sipper tube openings for open field box |  |  |  | 1 |  |  |  |  | Panel consists of 2 machined pieces of PVC sheet, with attached aluminum food grate, 3D printed support pieces for grate, metal screws to attach aluminum to 3D printed part(s), plastic screws and nuts. | JAX
4 | open field | Aluminum grate |  |  |  | 1 | [5 empty] | JAX
5 | open field | Food hopper to attach to grated panel |  |  |  | 1 |  |  |  | 3D printed |  | JAX
6 | open field | Lid for food hopper to attach to grated panel |  |  |  | 1 |  |  |  | 3D printed |  | JAX
7 | open field | Screw stops to hold panel to box (min 3-4 per box, 2 styles) |  |  |  | 4 |  |  |  | 3D printed plus plastic screws and nuts | Two styles - One to apply to flat face of panel and one to insert into sipper tube hole which is not being used. Our typical experiments use one water bottle, so we use 2 sipper tube screw stops to hold the bottom of the panel and 1 or 2 flat screw stops to hold the top. | JAX
8 | open field | [10 empty] | JAX
9 | open field | Box materials list from Jarek 8/6/19: | [9 empty] | JAX
10 | open field | white PVC sheet for box floor | Celtec® Expanded PVC Sheet, Celtec® 700, White, (6 mm x 48 in x 96 in), Satin / Smooth, Digital Print Grade | W01-00966 | Curbell Plastics | ? | 48"x96" sheet | 72.93 |  | https://www.curbellplastics.com/Shop-Materials/Product/Celtec-Expanded-PVC-Sheet/W01-00966/Celtec-Expanded-PVC-Sheet-Celtec-700-White-(6-mm-x-48-in-x-96-in)-Satin-Smooth-Digital-Print-Grade | White for box floor. Have them cut to 48" x 48" sheets before shipping | JAX
11 | open field | gray PVC sheet for box walls, door and panels | Celtec® Expanded PVC Sheet, Celtec® 700, Gray, (6 mm x 48 in x 96 in), Satin / Smooth, Digital Print Grade | W01-01004 | Curbell Plastics | ? | 48"x96" sheet | 80.22 |  | https://www.curbellplastics.com/Shop-Materials/Product/Celtec-Expanded-PVC-Sheet/W01-01004/Celtec-Expanded-PVC-Sheet-Celtec-700-Gray-(6-mm-x-48-in-x-96-in)-Satin-Smooth-Digital-Print-Grade | Gray for box walls. Have them cut to 48" x 48" sheets before shipping | JAX
12 | open field | corner brackets | Multipurpose 6061 Aluminum 90 Degree Angle with Round Edge, 1/8" Thickness, 1" High x 1" Wide Outside | 8982K4 | McMaster-Carr | 2? | 4ft pc | 10.55 |  | https://www.mcmaster.com/8982k4-8982K404 |  | JAX
13 | open field | door hinge | Steel Piano Hinge without Holes 1-1/2" Overall Width, 1/2" Long x 0.174" Diameter Knuckle | 15665A909 | McMaster-Carr | 1 | 3ft pc | 4.17 |  | https://www.mcmaster.com/15665a909-15665A3 |  | JAX
14 | open field | Aluminum extrusions (track to hold front panel) | T-Slotted Framing, Single Low-Profile Rail, 1" High x 1/2" Wide, Hollow | 47065T68 | McMaster-Carr | 2 | 2ft rail | 7.98 |  | https://www.mcmaster.com/47065t68-47065T681 |  | JAX
15 | open field | plastic screws | Nylon Pan Head Screws, Phillips, 1/4"-20 Thread, 1" Long | 94735A766 | McMaster-Carr | ? | pk of 100 | 9.01 |  | https://www.mcmaster.com/94735a766-93135A455 |  | JAX
16 | open field | plastic nuts | Nylon Hex Nut, 1/4"-20 Thread Size | 94812A700 | McMaster-Carr | ? | pk of 100 | 7.95 |  | https://www.mcmaster.com/94812a700-94812A116 |  | JAX
17 | open field | plastic nuts | Nylon Hex Nut, 10-32 Thread Size | 94812A600 | McMaster-Carr | ? | pk of 100 | 7.55 |  | https://www.mcmaster.com/94812a600-94812A415 |  | JAX
18 | open field | need a second size of plastic screws? | [9 empty] | JAX
19 | open field | 3D printed parts needed for assembled box: | [9 empty] | JAX
20 | open field | door latches (2) | [9 empty] | JAX
21 | open field | food grate support | [9 empty] | JAX
22 | open field | food hopper | [9 empty] | JAX
23 | open field | food hopper lid | [9 empty] | JAX
24 | open field | water bottle holder attachment | [9 empty] | JAX
25 | open field | flat front stops to hold panel to box (2 pieces to each) | [9 empty] | JAX
26 | open field | stops with sipper tube plugs, to hold panel to box (2 pieces to each) | [9 empty] | JAX
27 | open field | screw know | Screw knob for attaching camera to L-bracket | order information? | [7 empty] | JAX
28 | open field | Holder for NIR filter | 3D printed Ninjaflex |  | Jarek Tarpszo? | 1 |  |  |  |  | 3D printed Ninjaflex | JAX
29 | open field | frame to support NIR lights | frame to support NIR lights | LIST PLASTIC SHEET ORDER INFO | [6 empty] | Plastic sheets cut to size with interior rectangular opening. Light strips are attached to frame by zip ties through holes in frame. NIR light assembly rests on lip of VIS LED light ring. | JAX
30 | open field | Plastic pieces to hold touchscreens to Metroshelves using #2 screws and nuts. | Plastic pieces to hold touchscreens to Metroshelves using #2 screws and nuts. | Fabricated by Jarek | [6 empty] | Same plastic as used for NIR lights, cut as strips with drilled holes to attach touchscreen to stand top shelf. | JAX
31 | Misc | Quatricide (concentrated) | Pharmacal Research Laboratories QUATRICIDE DETERGENT GAL 4/CS | NC9400628 | Fisher Scientific | 0 | case of 4 gallons | 98.92 |  | https://www.fishersci.com/shop/products/quatricide-detergent-gal-4-cs/nc9400628 | Dilute per instructions on bottle (2 oz = 59ml per gallon) and apply by spraying and wiping with paper towels. | optional
32 | Misc | Bleach | [8 empty] | dilute to 10% for cleaning | optional
33 | Misc | 1/2 gallon screw top plastic jugs for dilute Quatricide or Bleach | [9 empty] | optional
34 | Misc | spray bottles | [8 empty] | Spray sipper tubes, food grates, panel, arena floor. Spray or rinse with water or ethanol after. | optional
35 | Misc | chemical stickers for spray bottles | [9 empty] | optional
36 | Misc | dust pan | [8 empty] | Useful for removing bedding and debris from dirty arena. Narrow dustpan of ~6" works better than a wide one. | optional
37 | Misc | key lock box | Master Lock Lock Box 5401D Set Your Own Combination Wall Mount Key Safe, 3-1/4 in. Wide | 571642471 | Walmart | 1 per room? |  | 21.09 |  | https://www.walmart.com/ip/Master-Lock-Lock-Box-5401D-Set-Your-Own-Combination-Wall-Mount-Key-Safe-3-1-4-in-Wide/19870072?wmlspartner=wlpa&selectedSellerId=0&adid=22222222227015875615&wl0=&wl1=g&wl2=c&wl3=40873682192&wl4=pla-60820287566&wl5=1018765&wl6=&wl7=&wl8=&wl9=pla&wl10=8175035&wl11=online&wl12=19870072&veh=sem&gclid=EAIaIQobChMI69PhhN6H4wIVhYKzCh2_3QcREAQYBSABEgKs4_D_BwE | If you want arenas in a locked room to limit access, mounting key lock box nearby is useful. | optional
38 | Misc | bin | Rubbermaid 3351 Bus Utility Box, 7-5/8 Gallon Capacity, Gray | RCP 3351 GRA | Rubbermaid | up to 1 per stand |  | 17.36 |  | https://www.cleanitsupply.com/p-1556/rubbermaid-3351-bus-utility-box-7-5-8-gallon-capacity-gray-rcp-3351-gra.aspx | For storing miscellaneous supplies on bottom shelf of stand | optional
39 | Misc | to fur mark light mice | Sharpie® Permanent Marker, Fine Point, Black | SAN30001 | W.B.Mason | 1 for many |  | 1.19 |  | https://www.wbmason.com/ProductDetail.aspx?ItemDesc=Sharpie-Permanent-Marker-Fine-Point-Black&ItemID=SAN30001&uom=EA&OCID=SAG&CID=NOV0920_PDP_OFFC_SHOPPINGOFFICESUPPLIES~OFFICESUPPLIES_GPA_OTRT&gclid=CjwKCAjwruSHBhAtEiwA_qCppsxYUZU_cZQmvyjlOHLoZnmEhc0nCKaLNMzazvcwOrgLXz6BnFZRABoC-f4QAvD_BwE | Visible in IR for a few days, how long depends on how much mice clean their fur. Apply immediately before placing mice in arena. | optional
40 | Misc | to fur mark dark mice | Hydrogen peroxide 30% stabilized ACS, VWR Chemicals BDH® | BDH7690-1 | VWR | 1 for many |  | 113.38 |  | https://us.vwr.com/store/product/10073707/hydrogen-peroxide-30-stabilized-acs-vwr-chemicals-bdh | Apply with cotton swabs 3-4 days before placing mice in arena, so they have time to return to 'natural' behavior prior to expt start. Marks persist at least a few weeks. | optional
41 | Misc | cotton swabs |  |  |  |  | pack of 100 |  |  |  | to apply hydrogen peroxide to mouse fur | optional
42 | Misc | 15ml Falcon tubes |  |  |  |  | pack of 50 |  |  |  | to hold hydrogen peroxide for soaking cotton swabs | optional
43 | Misc | scale to weigh mice, grain, water (need to 700g for food and water) | [9 empty] | optional
44 | Misc | camera to remotely view externally mounted water bottles | Wyze Cam Pan 1080p Pan/Tilt/Zoom Wi-Fi Indoor Smart Home Camera with Night Vision, 2-Way Audio, Works with Alexa & the Google Assistant, White - WYZECP1 | WYZECP1 | Amazon | 1 for several grouped together | ea | 47.2 |  | https://www.amazon.com/Wyze-1080p-Indoor-Camera-Vision/dp/B07DGR98VQ/ref=sr_1_15?dchild=1&keywords=camera+video+wifi+pan+tilt&qid=1615073201&s=electronics&sr=1-15 | Wireless, sees in infrared as well as visible, livestream accessed via iPhone. | optional
45 | Misc | camera to remotely view externally mounted water bottles | AXIS M5054 PTZ Network Camera | M5054 | Axis | 1 for several grouped together | ea | 500 |  | https://www.axis.com/products/axis-m5054 | Ethernet, sees in visible light only, managed by institute IT, requires site license, appropriate if IT security is a concern | optional
46 | Stand | shelves | Metro A2424NK3 Super Adjustable Metroseal 3 Wire Shelf - 24" x 24" | A2424NK3 | WebstaurantStore | 3 | ea | 55.5 | 166.5 | https://www.webstaurantstore.com/metro-a2424nk3-super-adjustable-metroseal-3-wire-shelf-24-x-24/461A2424NK3.html | includes shelf to pole connectors | Collaborator Lab
47 | Stand | fixed poles | Metro 74PK3 Stationary Super Erecta SiteSelect 74" Metroseal 3 Post | 46174PK3 | WebstaurantStore | 4 | ea | 16.5 | 66 | https://www.webstaurantstore.com/metro-74pk3-stationary-super-erecta-siteselect-74-metroseal-3-post/46174PK3.html | Use fixed poles OR poles with wheels for rolling stand | Collaborator Lab
48 | Stand | poles for rolling stand | Metro 74UPK3 Mobile Super Erecta SiteSelect 74" Metroseal 3 Post | 46174UPK3 | WebstaurantStore | 4 | ea | 16.5 | 66 | https://www.webstaurantstore.com/metro-74upk3-mobile-super-erecta-siteselect-74-metroseal-3-post/46174UPK3.html | Use fixed poles OR poles with wheels for rolling stand | Collaborator Lab
49 | Stand | wheels for rolling stand | Regency Replacement 5" Polyurethane Shelving Stem Caster with Brake | 4605MPBX | WebstaurantStore | 4 | ea | 8.99 | 35.96 | https://www.webstaurantstore.com/regency-replacement-5-polyurethane-shelving-stem-caster-with-brake/4605MPBX.html | Use fixed poles OR poles with wheels for rolling stand | Collaborator Lab
50 | open field | water bottle holder | Lixit Glass Bird Water Bottle 16 oz 5/16 in Tube GB16S | GB16S | Amazon | 3 | ea | 23.99 | 71.97 | https://www.amazon.com/Lixit-Glass-Water-Bottle-GB16S/dp/B0006B7FMO/ref=sr_1_15?crid=3RQAD4QQ3YJ65&keywords=lixit+water+bottle&qid=1559739705&s=gateway&sprefix=lixit+wa%2Caps%2C140&sr=8-15 | Each Lixit unit contains 1 water bottle holder, 1 glass water bottle, 1 rubber stopper and 1 sipper tube with ball. We modify the holder, replace glass bottles with standard JAX plastic bottles (JAX has water bottles with 2 size mouths. We use only the larger mouth bottles, which match the Lixit opening size. Metal holder has to have extra hole drilled, top cut off to clear door and 3D printed track attachment added. | Collaborator Lab
51 | open field | sipper tube | VWR Bent Sipper Tube, 12.7 cm (5") Long | 10718-064 | VWR |  | ea | 23.6 |  | https://us.vwr.com/store/product/10718-330/sipper-tube-econo-25-strght-e1027 | tube length to be cut to match Lixit tube, then replace Lixit tube in stopper. Can be used uncut, but that would mean the sipper tube sticks up in the bottle beyond the stopper, making some of the water unusable. | Collaborator Lab
52 | open field | rubber stoppers | #7 or #6 rubber stoppers |  | Lyon Scientific |  | lb (sold by wt?) | 22.83 |  | https://www.lyonscientific.com/products/rubber-stoppers?variant=5248548405285&currency=USD&utm_campaign=gs-2019-04-24&utm_source=google&utm_medium=smart_campaign | 2nd stopper source since we want 2x number of stands (#7 rubber stopper with 1 hole). Also purchase #6 stoppers to use with 50ml conical tubes for experiments with less water (like drugs in water). | Collaborator Lab
53 | open field | Flanges with set screws | HJ Garden 2pcs 14mm Flange Shaft Coupling High Hardness Metal Flanged Joint Guide Shaft Support Coupler for DIY Model Shaft Connection |  | Amazon Prime | 1 | 2 flanges | 12.69 |  | https://www.amazon.com/HJ-Garden-Coupling-Hardness-Connection/dp/B07MNT9VXW/ref=sr_1_1_sspa?keywords=Flange+Shaft+Coupling+14mm&qid=1556732320&s=industrial&sr=1-1-spons&psc=1 | Includes Allen wrench, for attachment of camera and lights to stand. | Collaborator Lab
54 | open field | screws to attach flange pairs to shelves | Steel Pan Head Phillips Screws, 2-56 thread, 7/8" long |  | McMaster-Carr | 3 screws | pk of 100 | 3.95 |  | https://www.mcmaster.com/90272A085 | For attachment of camera and lights to stand. | Collaborator Lab
55 | open field | nuts to attach flange pairs to shelves | 18-8 Stainless Steel Hex Nut, 2-56 thread size |  | McMaster-Carr | 3 nuts | pk of 100 | 3.11 |  | https://www.mcmaster.com/91841A003 | For attachment of camera and lights to stand. | Collaborator Lab
56 | open field | aluminum rod | 1/2" aluminum rods (~6" long) | [7 empty] | For attachment of camera and lights to stand. Purchase as longer pieces and cut to size. | JAX | need catalog number
57 | Camera or Lights | plastic sleeve for camera mount | 1/2 in. I.D. x 5/8 in. O.D. x 10 ft. Clear Vinyl Tubing |  | Home Depot | 1"lengh | sold in 10' length | 5.18 |  | https://www.homedepot.com/p/Everbilt-1-2-in-I-D-x-5-8-in-O-D-x-10-ft-Clear-Vinyl-Tubing-T10006010/304185142?MERCH=REC-_-searchViewed-_-NA-_-304185142-_-N& | plastic sleeve to adapt alum rod to L-bracket (~1" long? 1/2" inside, X outside diameter?). Start with long piece and cut off sections. Brian, vinyl is probably a lot softer than what we've been using. Should it work anyway? Is 1/2" inside diameter correct? | JAX
58 | Camera or Lights | Visible light source | VIS Light Ring with L bracket | 18040100K2 | F&V | 1 |  | 209 | 209 | https://www.fvlighting.com/r-300-se-daylight-led-ring-light-w-l-bracket.html |  | Collaborator Lab
59 | Camera or Lights | Washers | [8 empty] | Washers to use with L-brackets | Collaborator Lab | need catalog number
60 | Camera or Lights | power supplies for lights | AC to DC power adapter to use with VIS and NIR lights | VA0950 | F&V | 1 |  | 30 | 30 | https://www.fvlighting.com/ac-power-adapter-compatible-with-r-300-f3-k4000-and-more.html | VIS Light Ring and NIR lights each use same type Power Supply | Collaborator Lab
61 | Camera or Lights | camera mount adapter | Tripod Mount ace - Camera Mount Adapter | 2000029679 | Basler | 1 |  | 10.9 |  | https://www.baslerweb.com/en/products/vision-components/additional-accessories/tripod-mount-ace/#tab=specs | Brian confirm? This was listed as backup. Primary no longer available? | Collaborator Lab
62 | Camera or Lights | camera | Basler ace acA1300-75gm Monochrome GigE Camera | 106755 |  | 1 |  | 443 |  | https://www.edmundoptics.com/p/basler-ace-aca1300-75gm-monochrome-gige-camera/3427/ |  | Collaborator Lab
63 | Camera or Lights | lens | Tamron 12VM412ASIR 1/2" 4-12mm F/1.2 Infrared Manual C-Mount Lens | TA41212ML | B&H | 1 |  | 89.95 |  | https://www.bhphotovideo.com/c/product/414195-REG/Tamron_12VM412ASIR_12VM412ASIR_1_2_4_12_F_1_2.html |  | Collaborator Lab
64 | Camera or Lights | NIR filter | Hoya IR-80 (800nm), 50.8mm Sq., 2.5mm Thick, Colored Glass Longpass Filter | 66-070 | Edmund Optics | 1 | ea | 76.5 | 76.5 | https://www.edmundoptics.com/p/ir-80-800nm-508mm-sq-longpass-filter/20602/ |  | Collaborator Lab
65 | Camera or Lights | NIR light strips (940nm) | LightingWill DC12V 5Meter/16.4ft 72W SMD5050 300LEDs IR InfraRed 940nm Tri-chip White PCB Flexible LED Strips 60LEDs 14.4W Per Meter, Non-waterproof |  | Amazon Prime | 1 | 150cm long x 10mm wide | 95.99 | 95.99 | https://www.amazon.com/LightingWill-InfraRed-Tri-chip-Flexible-Waterproof/dp/B01DM9BNYW | 500cm x 10mm (1 5m spool) is enough for how many stands? | JAX
66 | Camera or Lights | connectors from NIR strips to power supply | IR to Power Supply Connector (attached to NIR light strips and sent in as a single unit) | LightingWill 10pcs Pack Strip to DC female plug Solderless Snap Down 2Conductor LED Strip Connector for Quick Splitter Connection of 10mm Wide 5050 5630 Single Color Flex LED Strips | Amazon Prime | 1 | 10 pack | 9.99 |  | https://www.amazon.com/LightingWill-Solderless-Connector-Turning-Connection/dp/B01DM7HOAG/ | 1 pack of 10 | JAX
67 | Camera or Lights | connectors from NIR strip to NIR strip | IR to IR Connector (attached to NIR light strips and sent in as a single unit) | LightingWill 10pcs Pack Strip to Strip with Wire Solderless Snap Down 2Conductor LED Strip Connector for 10mm Wide 5050 5630 Single Color Flex LED Strips | Amazon Prime | 5 | 10 pack | 7.99 |  | https://www.amazon.com/LightingWill-Solderless-Connector-Turning-Connection/dp/B01DM7H7B2/ | 2 packs of 10 to have spares | JAX
68 | Camera or Lights | power supplies for lights | AC to DC power adapter to use with VIS and NIR lights | VA0950 | F&V | 1 |  | 30 | 30 | https://www.fvlighting.com/ac-power-adapter-compatible-with-r-300-f3-k4000-and-more.html | VIS Light Ring and NIR lights each use same type Power Supply | Collaborator Lab
69 | Camera or Lights | Powered USB Hub | Wavlink 4-Port USB 3.0 Hub with One 2.4A Fast Charging Port (BC1.2 / iPad / Cellphone), Super Speed USB 3.0 Hub Up To 5Gbps, With 12V/2A Power Adapter, LED Indicator, Hot Swapping, Plug and Play | 9SIA6PF5MP1967 | New Egg | 1 |  | 22.99 | 22.99 | https://www.newegg.com/wavlink-ws-uh3042p1-usb/p/0J2-008P-00034 |  | Collaborator Lab
70 | Camera or Lights | CAT7 Ethernet cable | AmazonBasics RJ45 Cat7 Network Ethernet Patch Internet Cable - 7 Feet |  | Amazon | 3 |  | 9.92 | 29.76 | https://www.amazon.com/AmazonBasics-Network-Ethernet-Patch-Cable/dp/B013PUMUOK/ |  | Collaborator Lab
71 | Camera or Lights | CAT6 Ethernet cable | AmazonBasics RJ45 Cat-6 Ethernet Patch Internet Cable - 14 Feet (4.3 Meters) |  | Amazon | 1 |  | 7.99 | 7.99 | https://www.amazon.com/dp/B00N2VJ2CG/ref=twister_B00O1RTQJE?_encoding=UTF8&psc=1 |  | Collaborator Lab
72 | Camera or Lights | PoE Injectors | TRENDnet Gigabit Power Over Ethernet Plus (PoE+) Injector,Converts Non-PoE Gigabit to PoE+ or PoE Gigabit, Network Distances up to 100 M (328 Ft.), TPE-115GI |  | Amazon | 1 |  | 29.99 | 29.99 | https://www.amazon.com/TRENDnet-Ethernet-Supported-Auto-MDIX-TPE-113GI/dp/B00BK4W8TQ?th=1 |  | Collaborator Lab
73 | Computer | External hard drive | 4TB hard drive | Toshiba Canvio Basics 4TB Portable External Hard Drive USB 3.0, Black (HDTB440XK3CA) | Amazon |  |  | 89.99 | 0 | https://www.amazon.com/Toshiba-HDTB330XK3CB-Canvio-Portable-External/dp/B07KQPKJXH/?th=1 |  | Collaborator Lab
74 | Computer | TX2 Computer | TX2 Computers | eBOX560-900-FL | Axiomtek | 1 |  | 1185 |  | https://www.axiomtek.com/Default.aspx?MenuId=Products&FunctionId=ProductView&ItemId=24544&upcat=144 | price not shown on website. Do we have a quote? Previous quote $1185. Get a new quote. | Collaborator Lab
75 | Computer | Nvidia computer | NVIDIA Jetson AGX Xavier Developer Kit | B07G5FCJQB | Amazon | 3 | ea | 709 |  | https://www.amazon.com/dp/B07G5FCJQB | Use TX2 OR Nvidia, not both | Collaborator Lab
76 | Computer | USB-C for Nvidia | Amazon Basics USB 3.1 Type-C to 3 Port USB Hub with Ethernet Adapter - Black | B01M7PL2WP | Amazon | 3 | ea | 21.95 |  | https://www.amazon.com/AmazonBasics-Type-C-Port-Ethernet-Adapter/dp/B01M7PL2WP/ | not used with TX2 | Collaborator Lab
77 | Computer | Touchscreens | GeeekPi 7 Inch 1024x600 Capacitive Touch Screen HDMI Monitor TFT LCD Display for Raspberry Pi/Beagle Bone Black/Windows 10 and MacBook Pro |  | Amazon | 1 |  | 59.99 |  | https://www.amazon.com/GeeekPi-1024x600-Capacitive-Monitor-Raspberry/dp/B075QCXLPF/ | Now only needed for on-site calibration or control. Instead user browser user interface? | Collaborator Lab
78 | Computer | APC | APC Back-UPS Pro 1500VA UPS Battery Backup & Surge Protector (BR1500G) |  | Amazon | 1 for 2 |  | 239.99 |  | https://www.amazon.com/dp/B003Y24DEU/ref=cm_sw_r_cp_ep_dp_jfrSAbME5T3P1 | Can use one unit to supply needs of two adjacent stands. | Collaborator Lab
79 | Computer | Power strip | Belkin Power Strip Surge Protector with 8 AC Multiple Outlets, 8 ft Long Flat Plug Heavy Duty Extension Cord for Home, Office, Travel, Computer Desktop, Laptop & Phone Charging Bricks (2,500 Joules) |  | Amazon | 1 for 2 |  | 19.99 |  | https://www.amazon.com/dp/B002QPZXH2/ref=cm_sw_r_cp_ep_dp_XFdIAb737RSX9 | Can use one unit to supply needs of two adjacent stands. | Collaborator Lab
80 | Camera or Lights | Timer | myTouchSmart Digital Timer, 2 Grounded Outlets, Simple Setup, Daily and 7-Day Settings, Weather Resistant, Ideal for Indoor, Outdoor, Seasonal, Christmas Lights, 33862, BLACK |  | Amazon | 1 for lots |  | 19.5 |  | https://www.amazon.com/dp/B01MZ2DNTS/ | Visible lights from many arenas can be plugged into 1-2 power strips then into 1 timer. | Collaborator Lab
81 | Camera or Lights | Extention cord | Tripp Lite Power Cord Extension Cable, Heavy Duty, 14AWG, 5-15P to 5-15R, 15A, 15' (P024-015), 15 ft. , Black |  | Amazon | 0 |  | 22.87 |  | https://www.amazon.com/Tripp-Lite-Power-Extension-P024-015/dp/B01ILE5ZUG/ref=asc_df_B01ILE5ZUG/?tag=hyprod-20&linkCode=df0&hvadid=167146065113&hvpos=&hvnetw=g&hvrand=15662396485706794251&hvpone=&hvptwo=&hvqmt=&hvdev=c&hvdvcmdl=&hvlocint=&hvlocphy=1018765&hvtargid=pla-306408711104&psc=1 | 1 from wall to APC when no nearby outlet, if needed. | Collaborator Lab